## Slide 1
Anatoli Zelenski, BNL
Polarization techniques.
Polarized proton (deutron) H-(D-) sources.
Polarized 3He++ ion sources.
Polarized heavy (radioactive) ion beams .
USPAS-2020
Polarized Ion Sources

## Slide 2
High intensity polarized H- ion source and “Siberian Snakes” made possible the high luminosity RHIC operation with colliding polarized protons beams to study:
proton spin structure,
fundamental tests of QCD and electro-weak interaction.

## Slide 3
N+ -N-
P =
N+ +N-
External magnetic field direction
Electrons and protons have intrinsic angular momentum - Spin: S=1/2 and magnetic moments aligned with spin.
According to quantum mechanics: electron (or proton) spin can have only two projections (ms =+1/2 or –1/2) on the axis defined by magnetic field direction.
N+/N- =3, P=50%
Electron, proton beam polarization

## Slide 4
Spin filtering:
1. Focusing(defocusing) of atomic beam by sextupole magnets. ABS
2. Selective quenching of metastable atomic states.
3. Cross-section difference for opposite spin direction. Antiprotons.
Optical pumping. (OPPIS, collinear pumping, magnito- optics traps).
Spin-transfer collisions with polarized atoms.
Polarization in nuclear collisions.
Evaporation of polarized solid material.
Polarization techniques

## Slide 5: Filtering
The n+3He  γ + 4He absorption cross section is strongly spin-dependent:
This provides an effective process both for polarizing, and for analyzing the degree of polarization of neutron beams.
polarized outgoing neutrons
unpolarized incoming neutrons
polarized 3He
Polarized 3He Neutron Spin Filters

## Slide 6
Dissociator
Separating
6-pole magnet
RF
transition
Proton
source
Optically-
Pumped cell
Sona
transition
ECR ionizer
Cs cell
Cs beam
Plasma ionizer
Na cell,8%
He cell,70%
Optically-
Pumped cell
Polarization by electron spin
Polarization
Transfer to
protons
Ionization
 H-
 H-
 H-
 H-
 H+
H+
Optical pumping
Optically-
Pumped cell
H0
Source
ABS
LDS
Proton polarization techniques
OPPIS

## Slide 7
Hydrogen atoms with electron polarization: mJ=+1/2 trajectories
RF transition, polarization transfer
from electrons to protons
RHIC beam crossing
Breit-Rabi polarimeter
Ne-Fe-B permanent magnet
sextupole - 1.7 T
gradient 5.7 T/cm
Magnet system design, NIM A256, (2006)
T.Wise, W.Haeberli, H.Kolster et al.
Spin filtering technique. Atomic beam sources
Notes: 7

## Slide 8: H -jet polarimeter
Record 12.6∙1016 atoms/s
Atomic Beam intensity.
Atomic beam velocity ~ 1800 m/s
H-jet thickness at the collision point-1.2 ∙1012 atoms /cm2
The distance from the nozzle
To collision point is 127 cm
FWHM=5.5 mm
Notes: 8

## Slide 9: H - jet polarimeter.
Dissociator
Notes: 9

## Slide 10
Polarizer
Scattering
Asymmetry due to spin-orbital interaction in nucleus.

## Slide 11: Antiproton polarization in the storage ring
Selective loss in the polarized hydrogen storage cell.
Antiproton
beam

## Slide 12: Nearly resonant charge-exchange reactions which can be used for production of polarized ions. A.Belov
H0 + D+  H+ + D 0
D0 + H+  D+ + H0  ~ 5 10-15 cm2
H0 + D  H + D 0
D0 + H  D + H 0  ~ 10-14 cm2
3He0  + 4He++  3He++ + 4He 0  ~ 510-16 cm2
In order to produce polarized negative hydrogen ions it was necessary to have deuterium plasma consisting mainly from D+ and D- ions because slow polarized H- ions can be easily destroyed in collisions with plasma electrons
Plasma injector producing deuterium plasma enriched by D- ions with surface-plasma converter has been developed at INR.

## Slide 13
Polarized source of H-/D- and H+/D+ ions with nearly resonant charge-exchange plasma ionizer. INR Moscow

## Slide 14
Plasma generator for resonant charge-exchange ionizer
In order to produce polarized negative hydrogen ions it was necessary to have deuterium plasma consisting mainly from D+ and D- ions because slow polarized H- ions can be easily destroyed in collisions with plasma electrons.
Plasma injector producing deuterium plasma enriched by D- ions with
surface-plasma converter has been developed at INR.
D+, D-, e

## Slide 15: Oscillograms of polarized H- ion current (vertical scale-1mA/div) and unpolarized D- ion current (10mA/div)
Peak H- ion current 4 mA,
(3 mA-200us)
Polarization 85-90%
Normalized emittance 2  mm mrad
Unpolarized D- ion current 60 mA (~20 mA/cm2)
Pulse duration (FWHM) 170 s
Rep. rate 5 Hz
Unpolarized D- ion current, 60 mA

## Slide 16: COSY/Julich polarized H-/D- ion source
H0 + Cs0  H- + Cs +
D0 + Cs0  D- + Cs +
Polarized beam current
up to 50 A
Pulse duration is 10 ms (injector - cyclotron)
Polarizaton of beam injected into COSY ring,
up to 91,4 % !
R. Gebel - PSTP2007
Cs+ (30 keV) + Cs => Cs0
30 keV

## Slide 17
R.Gebel

## Slide 18: Polarized deuteron source for Dubna NUCLOTRON accelerator
    New source with nearly resonant charge-exchange ionizer is developed
    Project goal:
      10 mA D+ (Н+)
      polarization ~ 90 % from nominal vector polarization + (-) 1 and tensor polarization + 1,-2
    Parts of CIPIOS source from IUCF will be used.
    Collaboration between JINR (Dubna) and INR Moscow.
    Storage cell is planned to be used to reduce background current of H2+ ions

## Slide 19: The polarized ions source for the JINR accelerator complex
The project assumes the design and construction of a universal high-intensity source of polarized deuterons (protons) using a charge-exchange plasma ionizer. The output D+↑(H+↑) current of the source is expected to be at a level of 10mA. The polarization will be up to 90% of the maximal vector (±1) for D+↑ (H+↑) and tensor (+1, -2) for D+↑ polarization. Realization of the project is carried out in close cooperation with INR of the RAS (Moscow).

## Slide 20
5S1/2
5P1/2
λ = 795 nm
Circularly Polarized light
MJ = -1/2
MJ = +1/2
Spontaneous
decays.
nLJ
~ 100% polarized Rb atoms!
1.56 eV
W. Happer, Rev. Mod. Phys. 44, 170 (1972)
relaxation
Optical pumping application is limited by available laser wavelenghes!?
Recent progress in laser technology is very promising for new applications.
Depopulation optical pumping
Notes: Need to add info about N2 admixture, and the relaxation transitions it causes

## Slide 21
Polarized light
Linear
polarization
Circular
polarization
We will see world differently
in polarized light.

## Slide 22
The light can be polarized by scattering: filtering, reflection, refraction.
The circular polarized light transfer an angular momentum, which can be used for electron and nuclei polarization in atoms.
Laser beams are always polarized.
Laser beam is a powerful source of angular momentum:
10 W (795 nm)  4•1019 photons/sec

## Slide 23
Optical pumping of alkali metals
2S1/2
5/2
3/2
F =5/2
2P1/2
3/2
s
+
8Li
I
= 2
Li 671 nm
Na 590
K 770
Rb 795
Cs 894
Fr 817
Typical polarization = 60 – 70%
381 MHz
mF
-5/2 -3/2 -1/2 1/2 3/2 5/2
Valence electron absorbs light g tranfers polarization to nucleus via hyperfine coupling

## Slide 24
Collinear Optical Pumping polarizer at TRIUMF
Notes: Contrast with ISOLDE

## Slide 25: Electron polarization transfer in charge-excgange, or spin-exchange collisions.
Electron polarization transfer in charge-exchange collisions:
(H+ + Rb  H0  + Rb+) ~ 10-14 cm2
or
Electron polarization transfer in spin-exchange collisions:
 (H0 + Rb  H0  + Rb) ~ 2 ·10-15 cm2

## Slide 26
Rb+
H+
H+
Proton
source
Laser-795 nm
Optical pumping
Rb: NL(Rb) ~1014 cm-2
Sona
transition
Ionizer
cell
H-
Laser beam is a primary source of angular momentum:
10 W (795 nm) 4•1019 h/sec 2 A, H0 equivalent intensity.
Supperconducting solenoid 25 кГс
1.5 kG field
Charge-exchange collisions: ~10-14 cm2
Na-jet ionizer cell:
NL(Na)~ 3•1015 cm-2
Electron to proton polarization transfer
ECR-29 GHz
Н+ source
~50 mA, 2.8keV
Rb+
Spin-transfer polarization in proton-Rb collisions
Notes: 26

## Slide 27: Pulsed OPPIS with the atomic hydrogen injector at INR, Moscow, 1982-1990. First generation.
Atomic H injector
Helium ionizer cell
~80% ionization
efficiency.
Lamb-shift polarimeter
H- current 0.4mA, P=65%

## Slide 28
Atomic H
injector
He-cell
H-
High-brightness proton beam inside strong 2.5 T solenoid field
produced by atomic H beam ionization in the He-gas ionizer cell
The primary proton beam and
Equivalent atomic H intensity
current is 2.5 A at 6.5 keV beam energy !
25 kG
1.5 kG
OPPIS with atomic H0 injector
H0

## Slide 29
Atomic H injector
Helium ionizer cell
~80% ionization
efficiency.
Lamb-shift polarimeter
H- current 0.4mA, P=65%
Pulsed OPPIS with the atomic hydrogen injector at INR, Moscow, 1982-1990. First generation.

## Slide 30
10 keV
80%
H0 + He → H+ + He + e
Atomic H0 ionization efficiency in the He cell

## Slide 31
H+
H0
H+
H0
H-
H2
He
Rb
Na
CP1
Na-jet
cell
Atomic H
injector
He-ionizer
cell
Neutralizer
H2-cell
TMP1
Rb-cell
Sona
transition
OPPIS with atomic H0 injector and He-ionizer cell at RHIC, 2012-19

## Slide 32
BNL, BINP, Novosibisk, INR, Moscow- collaboration
The RHIC OPPIS with atomic hydrogen injector, Run-2013

## Slide 33
Plasmatron
4-grid proton
extraction system
H2 Neutralization cell
~ 2.5 A
equivalent
H0 beam
FABS produces 100-150 mA equivalent H0 beam intensity
Within the Na-jet ionizer acceptance.
“Fast Atomic Beam Source”, FABS, BINP, Novosibirsk,2011

## Slide 34
2.5-5.0 A of
proton beam
At 6-10 keV
energy
1820 holes ,
1.0 mm in diameter
R1, grid1-180 cm
R2, grid2-150 cm
R3, grid3-120 cm
R4, grid4-120cm
5.0 cm
FABS 4-grid (spherical) Ion Optical System
with “geometrical focusing

## Slide 35
H2-neutralizer
cell
High-brightness
Proton source
H-
FC
H+
Bending
magnet
200 cm
SEM
Deflecting
magnet
Atomic H beam profile measurements vs. distance from the source
Movable CEM beam
Profile monitor

## Slide 36
Re-1.98 cm at the distance:
L=200 cm-from the source
Re-1.40 cm at the distance:
L=100 cm-from the source
FWHM-2.3 cm
FWHM-3.25 cm
Atomic beam intensity profile vs. distance from the source.

## Slide 37
Re
Distance from the source: L-cm
New IOS#2
IOS#1
Re- half-width of beam intensity profile at -1/e level.
Beam profile FWHM = 1.67 Re
Total equivalent H beam intensity 2.5A
Atomic beam intensity profile vs. distance from the source measured with the movable secondary emission monitor
~300 mA/cm2

## Slide 38
H+(70%)
H0(6.5 keV)
H0(2.5 keV)
-4.0 kV
-4.1 kV
-4.1 kV
H0(6.5keV)
He-ionizer cell
Rb -cell
H0(30%)
Deceleration
H0 + He → H+ + He + e
-2-3 kV
+0.1kV
Residual un-polarized H0 beam component suppression
by energy separation

## Slide 39
He-pulsed
valve
3-grid beam
deceleration
system
He-ionizer cell and 3-grid energy separation system

## Slide 40
IxB-pulsed valve operation principle
Lorentz (Laplace) force moves the flexible conducting plate in the high (~ 3-5 T) magnetic field.
For I=10 A, L=5 cm, F=2.5 N. Current pulse duration ~100-500 us
B
I
B
F
I

## Slide 41: Optical pumping of Rb vapor, 784.7 nm pulsed Cr:LISAF laser
Rb
h
Proton beam.
28 mm
Optical pumping by σ+: ΔmJ=1
Spontaneous radiation: ΔmJ=0, 1.
mJ -1/2 +1/2
Rb pol.
Profile.
Rb vapor cell.

## Slide 42
Transversal vapor flow in the Na-jet cell.
Reduces sodium vapor losses for 3-4 orders
of magnitude, which allow the cell aperture
increase up to 3.0 cm .
Nozzle
500deg.C
Collector
~150 deg.C
Boiler
~500 deg.C
NL ~2·1015 at/cm2
L ~ 2-3 cm
Sodium-jet ionizer cell
Return
line
Notes: 42

## Slide 43
Na-jet cell is isolated and biased to – 32 keV. The H- beam is accelerated in a two-stage acceleration system
-32 kV
H-
35 keV
39 keV
-28 keV
-15 keV
3 keV
7 kev
H0
H- beam acceleration to 35 keV at the exit of Na-jet ionizer cell

## Slide 44: H- beam acceleration to 35 keV at the exit of Na-jet ionizer cell.
Na-jet cell is isolated and biased to – 32 keV. The H- beam is accelerated in a two-stage acceleration system.
-32 kV
H-
35 keV
-28 keV
-15 keV

## Slide 45
Variable collimator
In DB2 to improve
Energy resolution.
Lamb Shift
Polarimeter
FC4
35.0 keV
FC-Tk1,
750keV
OPPIS
23.7 deg bender
RFQ
Low Energy Beam Transport line
Notes: 45

## Slide 46
Rb-cell
Temp.-105 deg.
H- ion beam current accelerated to 200 MeV, June-2017
1.05 mA
300 us

## Slide 47
Rb-cell thickness-NL X 1013 atoms/cm2 | 4.5 | 5.5 | 7.5 | 10.4
Linac Current, μA | 440 | 520 | 740 | 1050
Booster Input ×1011 H-/pulse | 10.0 | 12.0 | 14.0 | 17.1
Pol. %, at 200 MeV | 86 | 86 | 84.5 | 83
Reliable long-term ∙operation of the source was demonstrated.
Very high suppression of un-polarized beam component was demonstrated.
Small beam emittance (after collimation for energy separation) and high transmission to 200 MeV.
Beam intensity and polarization at 200 MeV

## Slide 48
86.7%
86.4%
200 μA 400 μs pulse at 200 MeV
~ 4.8∙1011 H-/pulse
Polarization measurement in 200 MeV polarimeter.

## Slide 49
OPPIS
LINAC
Booster
AGS
RHIC
(2.0-2.5) ∙1011 p/bunch, P-70%
1.0 mA x 300us→18∙1011 polarized H- /pulse
9.0 ∙1011 polarized H- /pulse at 200 MeV
routinely in Run-17, Polarization-85%
(2.5-3.0) ∙1011 protons /pulse at 2.3 GeV
~1.8∙1011 p/bunch, P~65-75% at 100 GeV
P ~ 60-65% at 255 GeV
RHIC Polarized beam in Run 2013-17
Notes: 49
49

## Slide 51
3He polarization inside the EBIS in high (5.0T) magnetic field. No polarization losses during beam transport through gradient magnetic field.
EBIS is used for efficient ionization and accumulation of polarized 3He++ ions to the full capacity 1012 total charge (5∙1011 , 3He++ ions).
Production of polarized 3He++ beam in EBIS. A.Zelenski, J.Alessi, ICFA Beam Dynamics
Newsletter 30, p.39, (2003)
Existed 3He++ sources intensity- 50nA, Birmingham
Notes: 51

## Slide 52
Radial trapping of ions by the space charge of the electron beam.
Axial trapping by applied electrostatic potentials at ends of trap.
The total charge of ions extracted per pulse is ~ (0.5 – 0.8) x ( number electrons in the trap ~1.0∙1012 )
Ion output per pulse is proportional to the trap length and electron current.
Ion charge state increases with increasing confinement time.
Output current pulse is independent of species or charge state!
Principle of EBIS Operation
10A, 10-20 keV
Notes: 52

## Slide 53: “Extended” EBIS upgrade with new “injector” solenoid for polarized 3He++ ion production
Optically-pumped 3He cell
3He+
Up to 2×1011 3He++ ions/pulse
Ionization to 3He++
Polarization and ionization in high magnetic field
will produce 3He++ ion beam with P ≥ 80%

## Slide 54
5.0 T field, about 1.0 T field at minimal solenoid separation-30 cm
Extended EBIS superconducting solenoids, April 2018

## Slide 55
Circular polarized
Laser light at 1083 nm

## Slide 56: 3He optical pumping in high 2.0-5.0T magnetic field

## Slide 57: RF-discharge in 3.0 T magnetic field. 3He-cell diameter-25mm
RF-discharge parameters strongly affect maximum polarization.
Optimization of the 3He -cell geometry (smaller diameter?) and electrodes for RF input should improve polarization.

## Slide 58
Probe laser
absorption polarimeter
Jan,2016, Sealed cell, OD-30 mm-Pol.- 88%, Field-2.0 T

## Slide 59: Polarization measurements
Isolation Valve (IV)
open
Isolation Valve (IV)
closed
Polarization
equilibrium
80.3%
65.1%
34.9%

## Slide 60: 3He-gas purification and filling system
Modified Cryo-pump for 3He
purification and storage

## Slide 61: 3He –optically-pumped cell in the high magnetic field
Drift tube -40 kV
Insulator
Pulsed valve
3He-cell
Long, small diameter drift tube works like a 3He storage cell, which reduces gas load to the EBIS vacuum system and increases polarization due to ionization localization in the high magnetic field region.

## Slide 62
The development of the 3He polarizing apparatuses, the spin-rotator,
and the nuclear polarimeter at the 3He++ ion beam energy 6.0 MeV
(in the high-energy beam transport line after the EBIS drift-tube Linac)
is funded by the DOE Research and Development Funds for
the Next Generation Nuclear Physics Accelerators Facilities.
3He-4He-scattering
3He++ spin rotator and polarimeter in the EBIS HEBT line at 6.0 MeV beam energy

## Slide 63
STAR
PHENIX
AGS, 24GeV
LINAC
BOOSTER, 2.5 GeV
Pol. H- ion source
Spin Rotators
Siberian Snakes
200 MeV polarimeter
RHIC pC “CNI” polarimeters
RHIC
Absolute H-jet
polarimeter
L max = 1.6  1032 s-1cm-2 50 < √s < 510 GeV
AGS pC “CNI” polarimeter
EBIS
Polarization facilities at RHIC
Notes: 63

## Slide 64
The RHIC high intensity polarized H- source provides required beam intensity for present RHIC and future high-luminosity eRHIC collider operation.
The polarized 3He++ ion source on the basis of new EBIS injector is under development at BNL for future eRHIC collider.
High intensity un-polarized H- ion source development is in progress as a part of Linac intensity upgrade.
Summary
Notes: 64

## Slide 65
High-intensity un-polarized H- sources
And injector development at BNL

## Slide 66
OPPIS
Magnetron- S1
Magnetron -S2
To RFQ
Pulsed bending magnet, 45 deg.
Laser beam
23.7 deg. dc bending magnet
35 keV LEBT upgrade with three sources for Run-2020
Prototype of H- injector for high-energy accelerators, with high cost downtime

## Slide 67
Power efficiency : ~ 100 mA H-/1.5 kW
arc-discharge power
~0.6 sccm
BNL magnetron H- ion source

## Slide 68
L5-current after RFQ
750 keV energy
Current -80 mA
L4 –current
At injection
To RFQ-110 mA
Dec.5,2017, Arc-18A, Source current L1-130 mA, current after RFQ L5-80 mA
L1-130 mA
800 us pulse duration

## Slide 69
Polarization phenomena in nuclear physics at low , intermediate energies. IUCF, RCNP, COSY
All “meson factories” PSI, LAMPF, TRIUMF, INR Moscow had polarized proton beams.
Polarized beams in RHIC.
All electron-positron colliders operate with the polarized electron, positron beams.
Proposals.
Polarized beam at SSC.
Polarized beam at TEVATRON (FNAL), 1000GeV
Polarized beam at HERA, 920 GeV
Polarized beam at LHC.
Polarized beam at J-Park, 50 GeV